採購規格需求書

採購案名(中文)：量子位元微波控制系統一套

採購案名(英文)：Quantum bit microwave control system

履約期限： 決標次日起120日曆天

付款方式： 驗收合格後一次付款

保固期限： 自驗收合格之日起 保固1年

交貨地點： 中央研究院關鍵議題研究中心

| 系統規格General requirements of the system |  |
| --- | --- |
|  | The system should be based on an FPGA hardware, firmware, and software designed for pulsed quantum control and measurement. |
|  | The system should allow the generation of arbitrary pulse sequences on the FPGA with varying amplitudes, frequencies, phases, or durations, based on real-time instructions, without loading distinct sequences from software. |
|  | The system should include stable and low-noise DC channels for voltage biasing of gate electrodes. |
|  | Both FPGA based sequences and voltage biases should be controlled from the same software interface. |
|  | The system should allow delays up to 100 seconds between pulses for characterization of systems with long coherence/relaxation times |
|  | The system should allow real-time phase switching for advanced CPMG pulse cycling routines |
|  | The system should allow conditioning or modifying pulse generation, including pulse parameters noted in point 2, based on measurement results, realizing feedback latency less than 300ns |
|  | The system should include an extensive mathematical library to perform real-time ultrafast execution of arbitrary mathematical functions (e.g.: Bayesian estimates) using measurement data |

| 品項與系統規格 | 數量 |
| --- | --- |
| 量子位元微波控制系統(Quantum bit microwave control system) - 量子位元微波控制機架模組*6 • At least 2 power supplies, 6 fans • Maximum 8 FEMs per chassis • Maximum up to 64 channels (8 per module) • Waveform sampling rate> 1 GSa/s • Waveform vertical resolution > 14 bits • Jiter: < 0.4 ps - 量子位元微波控制微波訊號模組*6 • At least 8 outputs channels, and frequency 50 MHz – 10.5 GHz • Output channels resolution > 14 bits • Output channels sampling rate > 1 GSa • Output channels bandwidth > 700 MHz • At least 2 inputs channels, and frequency 2 GHz - 10.5 GHz •Input channels resolution > 10 bits •Input channels sampling rate > 2 GSa •Input channels bandwidth > 700 MHz •At least 8 digital channels •At least 16-core FPGA-based pulse processing unit - 量子位元微波控制低頻訊號模組*11 •At least 8 outputs channels, and frequency DC – 750 MHz •Output channels bandwidth > 300 MHz •At least 2 inputs channels, and frequency DC – 750 MHz •At least 8 digital channels •Rise time < 850 ps | 1套 |